Craig Dukes

24 August 2008

Wiener LV Power Supplies

Comments

We originally (October 11, 2006) received quotes for a PL508 supply with 3 MEH 2-7V pods and 3 MEH 15-30V pods.  We received a quote a year later for a PL506 with the same pods.  The PL506 has the advanced control and programming features of the PL512.

Power modules

- MEH:  individual outputs

- MDH:  dual outputs

- Modules cannot be easily replaced.  They need to be calibrated.  Frame stays, power supplies goes. Output power (MEH modules):

- 2-7 V MEH module

  - Max. current: 115 A.

  - Max. power:  550 W.

- 12-30 V MEH module

  - Max. current: 23 A.

  - Max. power:  550 W.

- Power module efficiency (MEH):  	5.0V/80A:  	79.4%

4.0V/80A:	~77%

3.3V/80A:	73.7%

24 V:		87.7%

- Regulation:

  - Static: 	< 25 mV @ ±100% load.

  - Dynamic:  	< 100 mV @ ±25% load.

- Ripple:  < 3 mV peak-to-peak

- Current limiting:  100% nominal value.

- Overvoltage protection:  125% of nominal value.

- Remote sense:

  - if disconnected supply will ramp up.

  - Two types:

    - Fast:  < 10 m run

    - Slow: < 100 m run

    - Need to specify which one we want

- Swivel nut to connect to terminals.

- Float isolation ±10V.

PL506 System

- Size:

  - 3U chassis

- Weight:  ? kg

- 6 single-output power modules:  6 channels

- Input:

  - Maximum current: 16 A

  - Maximum voltage:  264 VAC (240 + 10%)

  - Power factor:  0.997 (@230 VAC, 16A)

  - Power:  3840 W (240 VAC, 16A)

- Mains input/primary switcher efficiency:  95%

- Power:

  - Max. power @208 VAC:  2340 W (=208 V·16 A·0.74·0.95) (÷6=390 W)

  - Max. power @230 VAC:  2587 W (=230 V·16 A·0.74·0.95) (÷6=431 W)

  - Max. power @240 VAC:  2700 W (=240 V·16 A·0.74·0.95) (÷6=450 W)

- Identical modules to PL508

- Easier to swap supplies.

- Superior control features to PL508.  Can program different ramp rates, for example.

- Slightly more expensive

- No manual, identical to PL512

PL512 System

- 3U chassis:  same size as PL506

- 6 dual-output modules:  12 channels

- Maximum current: 16 A

- Maximum voltage:  264 VAC (240 + 10%)

- Used by STAR TOF

- Only uses MDH modules.

- Water cooling option.

- Max. power 2760W (= 230V∙16A∙0.75)

PL508 System

- First generation system

- Size:  434 mm W x 132 mm (3U) H x 355 mm D.  1U additional space needed for cooling, making height=178 mm.

- Weight:  ? kg

- 0-50°C operating temperature..

- 3U chassis provides 5 power modules + 1 16A AC input

- 6U chassis provides 10 power modules + 2 16A AC inputs

- Power factor:  0.997 (@230 VAC, 16A)

- Input power:

  - 16 A

  - 3 pin Hirschmann connector

- Up to 8 floating DC output channels

- Ramp rate fixed.

- Trip limits per channel; one shuts down entire supply.

- Used by D0 with slow sensing:  John Foglesong x2739, Linda Bagby, x3100

Miscellaneous

- Other users:

  - Alan Bross bought Wiener supplies for MICE.  Gary Drake from Argonne bought the supplies for MINOS.

- Delivery time:  12 weeks a.r.o.

- Warrenty:

  - 2 years parts and labor

  - Repairs done at Springfield, OH, with a turn around time of 48 hours

  - For complicated repairs returned to Germany

- Repairs:

  - To exchange power modules one has to open the power supply, replace the module, and do a full voltage and current calibration.  This has to be done by an authorized technician, either at Springfield, OH, or on site.