Craig Dukes

24 August 2008

High Voltage Power Supply Specifications

1.  Required specifications

- 0 – 500V

- 10 mA per PDB minimum current

- Floating

- 100 mV voltage resolution

2.  CAEN power supplies

CAEN A1511B specifications

- 12 channels per card.

- Ground:  individual floating channels up to 500V.

- Voltage range:  0 – 500V.

- Voltage resolution: 100 mV.

- Ripple:  < 30 mV pp.

- Voltage set accuracy:  ±0.3% ± 0.5V.

- Over/under voltage protection:  ±1 V.

- Current output:  1 mA / 10 mA full scale (internal jumper switches between the two ranges.

- Overcurrent protection 100 mA/1 A resolution.

- Connector:  DB-37.

CAEN SY1527LC Mainframe Specifications

- Size:  19″ x 8U x 720 mm.  (6U boards, 2U fans).

- Weight:  24 kg (with 1 A1532 and 1 A1531 power unit).

- 0-40°C operating temperature.

- Input power:

  - 100-230 VAC (Three-phase upon request).

  - 50/60 Hz, 15 A (@230 V)

  - 3400 W

  - Three power supplies:  A1531 for crate power, A1532 for cards

  - 2250 W maximum output power.

- Capacity:  16 boards/crate.

- Cooling:  15 cm (3.37U, or 5.91″) needed from top of crate to any other object.

3.  ISEG Power Supplies

Users:

- ATLAS silicon pixel (EHQ F607n_405)

- ALICE TRD (EDS 20 025p _203)

- ALICE TPC (EDS 20 025p_204)

- ALICE photon spectrometer (EHS 8605p_156)

- LHCb (EHQ 8605p_156)

MPOD LV/HV Mainframe

- Different types :

- Size : 9U bin (6U boards)

- Cooling:  Built-in fans with front air inlet and rear outlet

  - Capacity:  10.5 HV/LV slots, Controller takes first half slot

- Input power:

  - 100-240 VAC

  - 50-60 Hz

  - 16 A

  - Hirschmann 3-pin connector

  - 3000 W

  - No main fuse:  external overcurrent circuit breaker needed

- Interface:

  - Ethernet (TCP/IP), CANbus, USB Combi-interface

  - Simple Network Management Protocol (SNMP) via TCP/IP,  Object Linking and Embedding (OLE) for Process Control (OPC) server

  - Interlock connector

- Different modules:

MPOD EC-2H Mainframe

ISEG EHQ F6 05x_106F

- 16 channels per card

- Size:  6U

- Ground:  floating

- Voltage range:  0 – 500 V

- Maximum current:  10 mA/channel

- Connector:  REDEL 51 pin

ISEG EHQ F605p156_F

- 16 channels per card

- Size:  6U

- Ground:  floating

- Voltage range:  0 – 500 V

- Maximum current:  15 mA/channel

- Connector:  REDEL 51 pin